Calibrating your MaxxFlow HTC

Menu #

1.5 Density

Set bulk density in g/l. (Possible range 1 to 3000 g/l).

g/l

708

Enter the value.

E  Exits without change

  Changes the value and returns to the previous menu.

1.7 Drop Height

Enter drop height. This sets fixed-velocity automatically.

mm

Enter the value.

200

E  Exits without change

  Changes the value and returns to the previous menu.

2.1.1  Zero  Point Calibration

Start zero adjustment (empty pipe and conveying

turned on with no material).

Push ok to start zero calibration.

Cancel with no.

2.1.3  Full Calibration

Full Calibration with MaxxFlow HTC 100% filled in

no-flow condition.

Push ok to start zero calibration.

Cancel with no.

lbs

1. Now run the product for a material weigh test.

2. Calculate the over/under %.

(Actual – Displayed)/ Actual x 100 = % error

Now, add or subtract this error to 1 to get your Initial K Factor.

Example:  Actual weight 5250lbs = 100% correct  (truck scale verified)

Displayed weight after run is 4580lbs.  (5250 – 4580)/5250 x 100 = 0.127 or +12.7% under

K Factor (I) = 1 + (+0.127) = 1.13

Now adjust the MaxxFlow to compensate the difference by using the K Factor.

(If you are 12.7% low then you need a 1.13 calibration factor to make the system add 13% to the weight).

2.2 Factor

Correction factor affects directly the density measurement.

1.13

0.01 to 9.99

Default 1.0

Enter the value.

E  Exits without change

  Changes the value and returns to the previous menu.

1. Now run the product for a 2nd material weigh test.

2. Calculate the over/under %.

Example: Actual weight before test 5250lbs = 100% correct weight

Displayed weight after run is 5110lbs.  = (5250 – 5110)/5250 x 100 = +2.67% under

New K Factor = Old K Factor + % error = 1.13 + 2.67%  = 1.16

Continue doing material weight tests until the K Factor is optimized.

Note: The first calibration sets the K factor.  The 2nd, 3rd, etc. calibrations compensate the original K factor.

2.2 Factor

Correction factor affects directly the density measurement.

1.16

0.01 to 9.99

Default 1.0

Enter the value.

E  Exits without change

  Changes the value and returns to the previous menu.

1. Now run the product for a 3rd material weigh test.

2. Calculate the over/under %.

Example: Actual weight before test 5250lbs = 100% correct weight

Displayed weight after run is 5288lbs.  = (5250 – 5288)/5250 x 100 = -0.72% over

New K Factor = Old K Factor + % error = 1.16 + (-0.72%)  = 1.15

Now adjust the Maxx Flow to compensate the difference.

2.2 Factor

Correction factor affects directly the density measurement.

1.15

0.01 to 9.99

Default 1.0

Enter the value.

E  Exits without change

  Changes the value and returns to the previous menu.

Congratulations you have now calibrated your MaxxFlow HTC to your specific product.

If you need to calibrate the MaxxFlow HTC to a new product B, follow these procedures:

1. Understand that you DO NOT perform a Zero or Full Calibration!  Instead, you need to find the multiplier to apply so as to realize the correct Product B totalized weight.

2. Reset the MFE100 totalizer and run a material weight test on the new product at least 3 times.

2. Calculate the average % error to realize the Product B Multiplier.

RUN	Product B Actual                       Displayed Weight on MFE100

1	6570                                           5012

2	6540                                           5038

3	6595                                           5003

Avg	6568.33                                      5017.67

(Actual – Displayed)/ Displayed x 100 = % error,  (6568.33 – 5017.67)/ 5017.67 x 100 =  30.90%

Corrected Product B Weight = Displayed Weight + 30.90%

Calibration Worksheet

(Please fill in the blanks)

1.5 Density                     _______________g/l

1.7 Drop Height             _______________mm

2.1.1 Zero Point            _______________Range

_______________OFFSET

_______________Density

2.1.3 Full Calibration  _______________Density

Material Test (Product A)

Test1 	_____________Actual	____________Displayed

Calculate:  (Actual – Displayed)/ Actual  =  error    (error x 100 = error %)

Initial K Factor =  1 + error = _______________ K Factor (I)

Test2 	_____________Actual	____________Displayed 	_____________error % (2)

K Factor (2) = K Factor (I) + error % (2)

Test3	 ____________Actual 	____________Displayed 	_____________ error % (3)

K Factor (3) = K Factor (2) + error % (3)

Final K Factor (known as Calibration or Correction Factor)______________

Additional Product Material Tests (if more than 1 product)

Product B- Test1 	_____________Actual	____________Displayed

Product B- Test2 	_____________Actual	____________Displayed

Product B- Test3 	_____________Actual	____________Displayed

Product B- Average 	_____________Actual	____________Displayed

(Average Actual – Average Displayed)/ Average Displayed x 100 = _________% error

Corrected Product B Weight = Displayed Weight + % error =  __________________

Additional Product Material Tests (if more than 2 products)

Product C- Test1 	_____________Actual	____________Displayed

Product C- Test2 	_____________Actual	____________Displayed

Product C- Test3 	_____________Actual	____________Displayed

Product C- Average 	_____________Actual	____________Displayed

(Average Actual – Average Displayed)/ Average Displayed x 100 = _________% error

Corrected Product C Weight = Displayed Weight + % error =  __________________

Additional Product Material Tests (if more than 3 products)

Product D- Test1 	_____________Actual	____________Displayed

Product D- Test2 	_____________Actual	____________Displayed

Product D- Test3 	_____________Actual	____________Displayed

Product D- Average 	_____________Actual	____________Displayed

(Average Actual – Average Displayed)/ Average Displayed x 100 = _________% error

Corrected Product D Weight = Displayed Weight + % error =  __________________

Additional Product Material Tests (if more than 4 products)

Product E- Test1 	_____________Actual	____________Displayed

Product E- Test2 	_____________Actual	____________Displayed

Product E- Test3 	_____________Actual	____________Displayed

Product E- Average 	_____________Actual	____________Displayed

(Average Actual – Average Displayed)/ Average Displayed x 100 = _________% error

Corrected Product E Weight = Displayed Weight + % error =  __________________

GTS, Inc.

PO Box 799, Shalimar, FL 32579  USA

Phone: 850-651-3388   Fax: 850-651-4777

Email: info@onthelevel.com  -  Website: www.onthelevel.com